## BATCH
University of Northern British Columbia
Northern Analytical Laboratory Services
QA/QC Report File (Processing)
BATCH | A
Analyst: | Erwin
NALS Director: | Dr. Hossein Kazemian
Date: | 2026-02-13
Client Name: | Multiple Clients
Supervisor: | n/a
Client Company Name: | n/a
Reference File: | 2025NALS02540_2621_2623_2627_Feb_13_2026_OES
Number of Samples: | Multiple Samples
Detection Limits File:
ICPOES (Raw Data from QA/QC from UNADJUSTED DATA) |  | Ag 328.068 nm | Ag 338.289 nm | Al 396.152 nm | As 188.980 nm | As 193.696 nm | B 249.678 nm | B 249.772 nm | Ba 493.408 nm | Ba-A 455.403 nm | Be 234.861 nm | Be 313.042 nm | Bi 222.821 nm | Bi 223.061 nm | Ca 317.933 nm | Ca 422.673 nm | Cd 214.439 nm | Cd 226.502 nm | Ce 418.659 nm | Ce 446.021 nm | Co 238.892 nm | Co 258.033 nm | Cr 206.158 nm | Cr 267.716 nm | Cu-A 327.395 nm | Dy 340.780 nm | Dy 353.171 nm | Er 337.275 nm | Er 349.910 nm | Eu 397.197 nm | Eu 420.504 nm | Fe 234.350 nm | Fe 238.204 nm | Ga 294.363 nm | Ga 417.204 nm | Gd 335.048 nm | Gd 342.246 nm | Hf 263.872 nm | Hf 264.141 nm | Hg 184.887 nm | Hg 194.164 nm | Ho 339.895 nm | Ho 345.600 nm | K 766.491 nm | La 333.749 nm | La 379.477 nm | Li 610.365 nm | Li 670.783 nm | Lu 261.541 nm | Lu 291.139 nm | Mg 280.270 nm | Mg 285.213 nm | Mn 257.610 nm | Mn 294.921 nm | Mo 202.032 nm | Mo 204.598 nm | Na 588.995 nm | Na 589.592 nm | Nb 309.417 nm | Nb 313.078 nm | Nd 401.224 nm | Nd 406.108 nm | Ni 216.555 nm | Ni 231.604 nm | P 177.434 nm | P 178.222 nm | P 213.618 nm | Pb 220.353 nm | Rb 780.026 nm | S 180.669 nm | S 181.972 nm | Sb 206.834 nm | Sb 217.582 nm | Sc 335.372 nm | Sc 361.383 nm | Se 196.026 nm | Se 203.985 nm | Si 251.611 nm | Si 288.158 nm | Sm 359.259 nm | Sm 360.949 nm | Sn 189.925 nm | Sn 283.998 nm | Sr-R 407.771 nm | Sr-R 421.552 nm | Ta 263.558 nm | Ta 268.517 nm | Te 182.153 nm | Te 214.282 nm | Th 269.242 nm | Th 283.730 nm | Ti 334.941 nm | Ti 336.122 nm | Tl 190.794 nm | Tl 276.789 nm | Tl 351.923 nm | Tm 313.125 nm | Tm 342.508 nm | U 367.007 nm | U 385.957 nm | V 292.401 nm | V 309.310 nm | W 207.912 nm | W 209.475 nm | Y 360.074 nm | Y 371.029 nm | Yb 328.937 nm | Yb 369.419 nm | Zn 202.548 nm | Zn 213.857 nm | Zr 339.198 nm | Zr 343.823 nm
QA/QC SAMPLES NAMES | Comment | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (µg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (µg/L) | Conc. (µg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L)
IPC Blank after Calibration (NALS Blank) | IPC Blank
IPC CCV5 after Calibration | CCV 500ppb
IPC CCV5 after Batch | CCV 500ppb
Digest Method Blank (LRB) | 2026NALS02621 A01
Digest LFB | 2026NALS02621 A02
SAMPLE, RPD Calculated | 2026NALS02621 A03
SAMPLE DUPLICATE, RPD Calculated | 2026NALS02621 A04
LFM Background (10X Dil. of Sample) | 2026NALS02621 A03 - Dil. 10X
LFM Spike | 2026NALS02621 A05
Instrument LFB | Lab Fortified Blank 1ppm Jan 20/26 ICV
Instrument Interference Check | Interference Check 1ppm Jan 20/26 ICV
Instrument ICV | ICV LVL 5 Jan 27/26
Digest Method Blank Comparison to Sample Matrix | DO NOT INPUT ANY DATA HERE
Digest Method Blank Comparison to 2.2X RL | DO NOT INPUT ANY DATA HERE
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
Reportable Limit (3xMDL) | Reportable Limit (No Dilution)
Reportable Limit (3xMDL) | Reportable Limit (No Dilution)
YELLOW IS PREFERRED WAVELENGTH
ICPOES (Reportable Numbers) *DO NOT ADJUST*
Sample Name | Comment
PERFORMANCE CHECK CALCULATIONS
ADJUSTMENT FOR <RL *DO NOT ADJUST*
Sample Name | Comment
PERFORMANCE CHECK CALCULATIONS
PERCENT RECOVERIES, RPD, or PASS / FAIL *DO NOT ADJUST*
Sample Name | Comment
*THIS TABLE HAS TWO ROWS REMOVED (IF EDITING, TAKE THIS INTO ACCOUNT)*

## SUMMARY TABLE
REPORT ICPOES QA/QC DATA
| Total Metals | Units | IPC CCV5 after Calibration | IPC CCV5 after Batch | Digest Method Blank (LRB) | Digest LFB | SAMPLE DUPLICATE, RPD Calculated | LFM Spike | Instrument LFB | Instrument Interference Check | Instrument ICV | Digest Method Blank Comparison to Sample Matrix | Digest Method Blank Comparison to 2.2X RL
Silver
Aluminum
Arsenic
Boron
Barium
Beryllium
Calcium
Cadmium
Cerium
Cobalt
Chromium
Copper
Dysprosium
Erbium
Europium
Iron
Gallium
Gadolinium
Hafnium
Mercury
Holmium
Potassium
Lanthanum
Lutetium
Magnesium
Manganese
Molybdenum
Sodium
Niobium
Neodymium
Nickel
Phosphorus
Lead
Rubidium
Sulfur
Antimony
Selenium
Silicon
Samarium
Tin
Strontium
Tantalum
Tellurium
Thorium
Titanium
Thallium
Thulium
Uranium
Vanadium
Tungsten
Ytterbium
Zinc
Zirconium
| Total Metals | Units
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## INSTRUCTIONS
INSTRUCTIONS
NALS_WTRICPOES_QC_XL_20250725_Rev1
ICPOES RESULTS
1. Input raw data from ICPOES run from client sample into GREEN CELLS of the "BATCH" tab. Copy and Paste using "123 Values" function.
2. Check "SUMMARY TABLE" tab to view QC Results for selected elements
3. If a fail occurs see if it falls within the comments section selection, as per SOP.
CURRENT DETECTION LIMITS FILES
ICPOES | 2024NALS02081_ICPOES_FullMetals_DL_RL_Nov_04_2024_ICPOES
CURRENT INSTRUMENT TEMPLATES
ICPOES | NALS_ICPOES_TEMPLATE_AUG_18_2023
ACRONYMS
IPC | Instrument Perfomance Check
LRB | Lab Reagent Blank
LFB | Lab Fortified Blank
IC | Interference Check
LFM | Lab Fortified Matrix
CCV | Continuous Calibration Verification
ICV | Independent Calibration Verification
PERCENT RECOVERY EQUATION (EPA 200.7)
Percent Recovery = (Cs - C / s) *100
Cs = Sample Concentration
C = Sample Background
s = true analyte concentration
RELATIVE PERCENT DIFFERENCE EQUATION
RPD = (ABS(Determined Value-True Value)/True Value)*100
RELATIVE PERCENT DIFFERENCE EQUATION FOR DUPLICATES
RPD = (abs(R1-R2)/((R1+R2)/2))*100
ICV and CCV PERCENT RECOVERY EQUATION
Percent Recovery = (Cs/s)*100
Cs = Sample Concentration
s = true analyte concentration

## CODES
DO NOT EDIT THIS PAGE
INSTRUMENT LIST
ICPOES
ICPMS
ICPQQQ
ICPOES CODES
o
####
Uncal
over-range
UNITS
µg/L
mg/L
%
CONDITIONAL FORMATTING
<0.0000
1000
20
INSTRUMENT REPORTABLE LIMIT TABLE |  | Ag 328.068 nm | Ag 338.289 nm | Al 396.152 nm | As 188.980 nm | As 193.696 nm | B 249.678 nm | B 249.772 nm | Ba 493.408 nm | Ba-A 455.403 nm | Be 234.861 nm | Be 313.042 nm | Bi 222.821 nm | Bi 223.061 nm | Ca 317.933 nm | Ca 422.673 nm | Cd 214.439 nm | Cd 226.502 nm | Ce 418.659 nm | Ce 446.021 nm | Co 238.892 nm | Co 258.033 nm | Cr 206.158 nm | Cr 267.716 nm | Cu-A 327.395 nm | Dy 340.780 nm | Dy 353.171 nm | Er 337.275 nm | Er 349.910 nm | Eu 397.197 nm | Eu 420.504 nm | Fe 234.350 nm | Fe 238.204 nm | Ga 294.363 nm | Ga 417.204 nm | Gd 335.048 nm | Gd 342.246 nm | Hf 263.872 nm | Hf 264.141 nm | Hg 184.887 nm | Hg 194.164 nm | Ho 339.895 nm | Ho 345.600 nm | K 766.491 nm | La 333.749 nm | La 379.477 nm | Li 610.365 nm | Li 670.783 nm | Lu 261.541 nm | Lu 291.139 nm | Mg 280.270 nm | Mg 285.213 nm | Mn 257.610 nm | Mn 294.921 nm | Mo 202.032 nm | Mo 204.598 nm | Na 588.995 nm | Na 589.592 nm | Nb 309.417 nm | Nb 313.078 nm | Nd 401.224 nm | Nd 406.108 nm | Ni 216.555 nm | Ni 231.604 nm | P 177.434 nm | P 178.222 nm | P 213.618 nm | Pb 220.353 nm | Rb 780.026 nm | S 180.669 nm | S 181.972 nm | Sb 206.834 nm | Sb 217.582 nm | Sc 335.372 nm | Sc 361.383 nm | Se 196.026 nm | Se 203.985 nm | Si 251.611 nm | Si 288.158 nm | Sm 359.259 nm | Sm 360.949 nm | Sn 189.925 nm | Sn 283.998 nm | Sr-R 407.771 nm | Sr-R 421.552 nm | Ta 263.558 nm | Ta 268.517 nm | Te 182.153 nm | Te 214.282 nm | Th 269.242 nm | Th 283.730 nm | Ti 334.941 nm | Ti 336.122 nm | Tl 190.794 nm | Tl 276.789 nm | Tl 351.923 nm | Tm 313.125 nm | Tm 342.508 nm | U 367.007 nm | U 385.957 nm | V 292.401 nm | V 309.310 nm | W 207.912 nm | W 209.475 nm | Y 360.074 nm | Y 371.029 nm | Yb 328.937 nm | Yb 369.419 nm | Zn 202.548 nm | Zn 213.857 nm | Zr 339.198 nm | Zr 343.823 nm
|  | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (µg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (µg/L) | Conc. (µg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L)
Reportable Limit (3xMDL) | Reportable Limit (No Dilution) | 0.007 | 0.009 | 0.012 | 0.013 | 0.014 | 0.0017 | 0.0011 | 0.0007 | 0.001 | 0.0005 | 0.0015 | n/a | n/a | 0.004 | 0.02 | 0.0007 | 0.0007 | 0.001 | 0.002 | 0.0009 | 0.0014 | 0.0032 | 0.0005 | 0.0007 | 0.002 | 0.001 | 0.0003 | 0.001 | 0.0005 | 0.0003 | 0.0067 | 0.0025 | 0.004 | 0.005 | 0.001 | 0.001 | 0.002 | 0.001 | 7 | 7 | 0.002 | 0.001 | 0.002 | 0.0004 | 0.0003 | n/a | n/a | 0.0004 | 0.001 | 0.0007 | 0.0041 | 0.0007 | 0.0035 | 0.0023 | 0.0022 | 0.0069 | 0.0011 | 0.01 | 0.06 | 0.003 | 0.002 | 0.0017 | 0.0009 | n/a | n/a | 0.011 | 0.0053 | 0.007 | 0.056 | 0.034 | 0.011 | 0.007 | n/a | n/a | 0.018 | 0.025 | 0.004 | 0.003 | 0.001 | 0.002 | 0.01 | 0.005 | 0.0003 | 0.0005 | 0.003 | 0.004 | 0.1 | 0.01 | 0.01 | 0.003 | 0.0005 | 0.0005 | 0.01 | 0.03 | 0.02 | 0.002 | 0.001 | 0.03 | 0.012 | 0.0008 | 0.003 | 0.005 | 0.01 | n/a | n/a | 0.0003 | 0.0004 | 0.001 | 0.0007 | 0.0006 | 0.0005
Reportable Limit (3xMDL) | Reportable Limit (No Dilution) | <0.007 | <0.009 | <0.012 | <0.013 | <0.014 | <0.0017 | <0.0011 | <0.0007 | <0.0010 | <0.0005 | <0.0015 | n/a | n/a | <0.004 | <0.02 | <0.0007 | <0.0007 | <0.001 | <0.002 | <0.0009 | <0.0014 | <0.0032 | <0.0005 | <0.0007 | <0.002 | <0.001 | <0.0003 | <0.001 | <0.0005 | <0.0003 | <0.0067 | <0.0025 | <0.004 | <0.005 | <0.001 | <0.001 | <0.002 | <0.001 | <7 | <7 | <0.002 | <0.001 | <0.002 | <0.0004 | <0.0003 | n/a | n/a | <0.0004 | <0.001 | <0.0007 | <0.0041 | <0.0007 | <0.0035 | <0.0023 | <0.0022 | <0.0069 | <0.0011 | <0.01 | <0.06 | <0.003 | <0.002 | <0.0017 | <0.0009 | n/a | n/a | <0.011 | <0.0053 | <0.007 | <0.056 | <0.034 | <0.011 | <0.007 | n/a | n/a | <0.018 | <0.025 | <0.004 | <0.003 | <0.001 | <0.002 | <0.010 | <0.005 | <0.0003 | <0.0005 | <0.003 | <0.004 | <0.1 | <0.01 | <0.01 | <0.003 | <0.0005 | <0.0005 | <0.01 | <0.03 | <0.02 | <0.002 | <0.001 | <0.03 | <0.012 | <0.0008 | <0.003 | <0.005 | <0.01 | n/a | n/a | <0.0003 | <0.0004 | <0.0010 | <0.0007 | <0.0006 | <0.0005
PERCENT RECOVERIES | PERCENT | LOW END, % | UPPER END, %
IPC Blank (NALS Blank) | <RL |  |  | <-- EPA 200.7, Section 9.3.4
IPC CCV5 after Calibration | 10 | 90 | 110 | <-- EPA 200.7, Section 9.3.4
IPC CCV5 after Batch | 10 | 90 | 110 | <-- EPA 200.7, Section 9.3.4
Digest Method Blank (LRB) | <RL or up to 10% of sample matrix or 2.2X RL | 10 | 2.2 | <-- EPA 200.7, Section 9.3.1
Digest LFB | 20 | 80 | 120 | <-- EPA 200.7, Section 9.3.2
SAMPLE, RPD Calculated |  |  |  | <-- Standard Method for the Examination of Water and Wastewater, Section 1020 B. Quality Control, and BCMOE SECTION C METALS C-9 Quality Control Table
SAMPLE DUPLICATE, RPD Calculated | 20 | 80 | 120 | <-- Standard Method for the Examination of Water and Wastewater, Section 1020 B. Quality Control, and BCMOE SECTION C METALS C-9 Quality Control Table
LFM Background |  |  |  | <-- Used for calculation for LFM
LFM Spike | 30 | 70 | 130 | <-- EPA 200.7, Section 9.3.5
Instrument LFB | 20 | 80 | 120 | <-- EPA 200.7, Section 9.3.2
Instrument Interference Check | 30 | 70 | 130 | <-- EPA 200.7, Section 9.3.5, while not stated, using LFM values
Instrument ICV | 5 | 95 | 105 | <-- Standard Method for the Examination of Water and Wastewater, Section 1020 B. Quality Control 1. 'First Paragraph'
| All Metals Except Hg (mg/L) | Hg (ug/L)
SPIKE AMOUNT (UNDILUTED) | 0.1 | 10
CCV5 Amount | 0.5 | 50
ICV Amount | 0.5 | 50

## TRACKED CHANGES
DATE | VERSION | AUTHOR | EDITS
2025-07-03 | NALS_WTRICPOES_QC_XL_20250703_Rev0 | Erwin Rehl | Entire excel workbook was created by Erwin
2025-07-25 | NALS_WTRICPOES_QC_XL_20250703_Rev0 | Erwin Rehl | Updated ICV and CCV5 equation to not use blank subtraction as they are not diluted into matrix solution after being made. Included equation in "Instructions Tab".

## BATCH VERIFICATION
University of Northern British Columbia
Northern Analytical Laboratory Services
QA/QC Report File (Processing)
BATCH | A
Analyst: | Erwin
NALS Director: | Dr. Hossein Kazemian
Date: | 2025-06-11
Client Name: | Multiple Clients
Supervisor: | n/a
Client Company Name: | n/a
Reference File: | 2025NALS02286_2290_2293_2296
Number of Samples: | 1+2+1+2
Detection Limits File:
ICPOES (Raw Data from QA/QC from UNADJUSTED DATA) |  | Ag 328.068 nm | Ag 338.289 nm | Al 396.152 nm | As 188.980 nm | As 193.696 nm | B 249.678 nm | B 249.772 nm | Ba 493.408 nm | Ba-A 455.403 nm | Be 234.861 nm | Be 313.042 nm | Bi 222.821 nm | Bi 223.061 nm | Ca 317.933 nm | Ca 422.673 nm | Cd 214.439 nm | Cd 226.502 nm | Ce 418.659 nm | Ce 446.021 nm | Co 238.892 nm | Co 258.033 nm | Cr 206.158 nm | Cr 267.716 nm | Cu-A 327.395 nm | Dy 340.780 nm | Dy 353.171 nm | Er 337.275 nm | Er 349.910 nm | Eu 397.197 nm | Eu 420.504 nm | Fe 234.350 nm | Fe 238.204 nm | Ga 294.363 nm | Ga 417.204 nm | Gd 335.048 nm | Gd 342.246 nm | Hf 263.872 nm | Hf 264.141 nm | Hg 184.887 nm | Hg 194.164 nm | Ho 339.895 nm | Ho 345.600 nm | K 766.491 nm | La 333.749 nm | La 379.477 nm | Li 610.365 nm | Li 670.783 nm | Lu 261.541 nm | Lu 291.139 nm | Mg 280.270 nm | Mg 285.213 nm | Mn 257.610 nm | Mn 294.921 nm | Mo 202.032 nm | Mo 204.598 nm | Na 588.995 nm | Na 589.592 nm | Nb 309.417 nm | Nb 313.078 nm | Nd 401.224 nm | Nd 406.108 nm | Ni 216.555 nm | Ni 231.604 nm | P 177.434 nm | P 178.222 nm | P 213.618 nm | Pb 220.353 nm | Rb 780.026 nm | S 180.669 nm | S 181.972 nm | Sb 206.834 nm | Sb 217.582 nm | Sc 335.372 nm | Sc 361.383 nm | Se 196.026 nm | Se 203.985 nm | Si 251.611 nm | Si 288.158 nm | Sm 359.259 nm | Sm 360.949 nm | Sn 189.925 nm | Sn 283.998 nm | Sr-R 407.771 nm | Sr-R 421.552 nm | Ta 263.558 nm | Ta 268.517 nm | Te 182.153 nm | Te 214.282 nm | Th 269.242 nm | Th 283.730 nm | Ti 334.941 nm | Ti 336.122 nm | Tl 190.794 nm | Tl 276.789 nm | Tl 351.923 nm | Tm 313.125 nm | Tm 342.508 nm | U 367.007 nm | U 385.957 nm | V 292.401 nm | V 309.310 nm | W 207.912 nm | W 209.475 nm | Y 360.074 nm | Y 371.029 nm | Yb 328.937 nm | Yb 369.419 nm | Zn 202.548 nm | Zn 213.857 nm | Zr 339.198 nm | Zr 343.823 nm
QA/QC SAMPLES NAMES | Comment | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (µg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (µg/L) | Conc. (µg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L) | Conc. (mg/L)
IPC Blank after Calibration (NALS Blank) | IPC Blank | 0.0017 | -0.0002 u | -0.0042 u | -0.0066 u | 0.0037 | 0.0001 u | 0.0002 | 0.0000 u | 0 | 0.0000 u | 0.0004 | Uncal | Uncal | -0.0068 u | -0.0114 u | 0.0001 | -0.0001 u | 0.0000 u | 0.0002 | -0.0001 u | 0.0001 u | -0.0021 u | 0.0001 | 0.0001 u | 0.0004 u | 0.0000 u | 0.0002 | 0.0000 u | 0.0001 | 0.0000 u | 0.0015 | 0.0016 | 0.0006 | 0.0006 u | 0.0001 u | -0.0001 u | 0.0001 u | 0.0003 | 1.2945 | -2.3549 u | -0.0003 u | -0.0001 u | 0.0001 u | 0.0000 u | 0.0000 u | Uncal | Uncal | 0.0000 u | 0.0001 u | 0.0000 u | -0.0006 u | 0.0002 | 0.0012 | 0.0008 | 0.0019 | -0.0025 u | 0.0001 u | 0.0008 | -0.0053 u | 0.0002 | 0.0003 u | 0.0004 | -0.0003 u | Uncal | -0.0089 u | 0.0051 | 0.0006 | 0.0009 u | 0.0000 u | -0.0032 u | 0.0020 u | 0.0043 | Uncal | Uncal | 0.0017 u | -0.0037 u | -0.0001 u | 0.0002 | 0.0001 | -0.0002 u | -0.0038 u | -0.0002 u | 0.0000 u | -0.0001 u | 0.0001 u | 0.0003 u | 0.0008 u | 0.0007 u | -0.0012 u | 0.0001 u | 0.0004 | 0.0003 | 0.0008 u | -0.0047 u | -0.0045 u | -0.0003 u | 0.0000 u | 0.0002 u | -0.0002 u | -0.0003 u | 0 | 0.003 | 0.0044 | Uncal | Uncal | 0 | 0.0000 u | 0.0000 u | -0.0003 u | 0.0000 u | 0.0000 u
IPC CCV5 after Calibration | CCV 500ppb | 0.4009 | 0.4004 | 0.5087 | 0.4977 | 0.5047 | 0.4992 | 0.5045 | 0.4975 | 0.5006 | 0.5044 | 0.515 | Uncal | Uncal | 0.4992 | 0.5017 | 0.4998 | 0.5076 | 0.5011 | 0.4977 | 0.4989 | 0.4996 | 0.5028 | 0.502 | 0.5 | 0.505 | 0.5032 | 0.5128 | 0.5049 | 0.5144 | 0.5103 | 0.5152 | 0.5153 | 0.4994 | 0.5002 | 0.5011 | 0.5029 | 0.5001 | 0.499 | 49.5219 | 47.9035 | 0.5034 | 0.5005 | 0.5123 | 0.5055 | 0.5015 | Uncal | Uncal | 0.5006 | 0.4996 | 0.5099 | 0.5056 | 0.5041 | 0.5022 | 0.5 | 0.5002 | 0.5015 | 0.5128 | 0.5004 | 0.4861 | 0.4974 | 0.5019 | 0.5032 | 0.5006 | Uncal | 0.5344 | 0.513 | 0.4979 | 0.4861 | 0.4846 | 0.4936 | 0.4875 | 0.4852 | Uncal | Uncal | 0.504 | 0.5092 | 0.4997 | 0.4975 | 0.5082 | 0.5073 | 0.5062 | 0.503 | 0.4999 | 0.4994 | 0.51 | 0.5094 | 0.4998 | 0.5057 | 0.4938 | 0.5162 | 0.499 | 0.5035 | 0.5024 | 0.4957 | 0.502 | 0.5567 | 0.503 | 0.4891 | 0.5032 | 0.506 | 0.5207 | 0.4976 | 0.4948 | Uncal | Uncal | 0.5042 | 0.507 | 0.4997 | 0.5003 | 0.5038 | 0.5025
IPC CCV5 after Batch | CCV 500ppb | 0.4 | 0.3967 | 0.5042 | 0.4965 | 0.506 | 0.5002 | 0.5054 | 0.4992 | 0.5018 | 0.5041 | 0.5166 | Uncal | Uncal | 0.4985 | 0.4975 | 0.5003 | 0.5073 | 0.5014 | 0.4986 | 0.4995 | 0.5008 | 0.5043 | 0.5015 | 0.5005 | 0.5048 | 0.5043 | 0.5105 | 0.5048 | 0.5137 | 0.5108 | 0.5024 | 0.5024 | 0.5004 | 0.5011 | 0.5003 | 0.5039 | 0.4994 | 0.4977 | 49.3751 | 49.0272 | 0.503 | 0.5002 | 0.5143 | 0.5032 | 0.5027 | Uncal | Uncal | 0.5014 | 0.4994 | 0.5095 | 0.5027 | 0.5034 | 0.5016 | 0.5016 | 0.5029 | 0.5091 | 0.516 | 0.4974 | 0.4851 | 0.4989 | 0.5019 | 0.5032 | 0.5003 | Uncal | 0.5082 | 0.5068 | 0.4969 | 0.4867 | 0.4997 | 0.5025 | 0.4916 | 0.4873 | Uncal | Uncal | 0.5107 | 0.5116 | 0.5012 | 0.4998 | 0.5109 | 0.5075 | 0.5065 | 0.503 | 0.504 | 0.5028 | 0.5092 | 0.5068 | 0.5268 | 0.5038 | 0.495 | 0.5158 | 0.5009 | 0.503 | 0.503 | 0.4949 | 0.4948 | 0.5588 | 0.5042 | 0.4918 | 0.5029 | 0.5063 | 0.521 | 0.4966 | 0.4944 | Uncal | Uncal | 0.505 | 0.5097 | 0.5007 | 0.5018 | 0.5049 | 0.5038
Digest Method Blank (LRB) | 2025NALS02293 A01 | 0.0004 | -0.0002 u | 0.0012 u | -0.0036 u | 0.0041 | 0.0006 | 0.0006 | -0.0003 u | 0.0000 u | 0.0000 u | 0.0004 | Uncal | Uncal | -0.0042 u | -0.0023 u | 0.0001 | -0.0001 u | -0.0001 u | 0.0001 | 0.0000 u | -0.0001 u | -0.0009 u | 0.0001 | 0.0000 u | 0.0002 u | 0.0000 u | 0.0000 u | 0 | 0.0000 u | 0.0000 u | -0.0001 u | -0.0001 u | 0.0005 | 0.0005 | 0.0002 u | -0.0001 u | -0.0001 u | 0.0001 u | 0.5189 u | -1.5927 u | -0.0003 u | -0.0002 u | -0.0005 u | -0.0001 u | 0.0000 u | Uncal | Uncal | 0.0000 u | 0.0001 u | 0.0003 | 0.0005 u | 0.0000 u | 0.0012 | -0.0001 u | 0.0007 | 0.0003 u | 0.0025 | -0.0003 u | -0.0105 u | 0.0000 u | 0.0004 | 0.0008 | -0.0002 u | Uncal | 0.0067 u | 0.0076 | 0.0014 | 0.0013 | 0.0028 u | 0.0071 | 0.0012 u | 0.0016 u | Uncal | Uncal | -0.0034 u | -0.0004 u | 0.0092 | 0.0094 | 0.0002 | -0.0002 u | -0.0017 u | -0.0002 u | 0.0000 u | 0.0000 u | -0.0001 u | 0.0003 u | -0.0144 u | -0.0006 u | -0.0013 u | 0.0004 | 0 | 0.0001 | 0.0002 u | -0.0024 u | -0.0052 u | -0.0001 u | -0.0002 u | -0.0019 u | 0.0004 u | -0.0002 u | 0.0000 u | 0.0009 | -0.0017 u | Uncal | Uncal | 0 | 0.0000 u | 0.0005 | 0.0004 | 0.0000 u | 0.0000 u
Digest LFB | 2025NALS02293 A02 | 0.0297 | 0.0231 | 0.1032 | 0.0967 | 0.0966 | 0.102 | 0.1024 | 0.1033 | 0.1043 | 0.1031 | 0.1058 | Uncal | Uncal | 0.0974 | 0.1 | 0.1021 | 0.1031 | 0.0924 | 0.0938 | 0.1038 | 0.1034 | 0.1015 | 0.103 | 0.1032 | 0.1015 | 0.1032 | 0.1026 | 0.1026 | 0.1054 | 0.1047 | 0.1032 | 0.1037 | 0.0983 | 0.1002 | 0.1016 | 0.1024 | 0.1 | 0.099 | 9.1142 | 7.971 | 0.1026 | 0.1025 | 0.1006 | 0.1027 | 0.1027 | Uncal | Uncal | 0.1024 | 0.1028 | 0.1044 | 0.1042 | 0.102 | 0.1026 | 0.105 | 0.1074 | 0.1037 | 0.1048 | 0.0851 | 0.0779 | 0.0977 | 0.103 | 0.1045 | 0.1029 | Uncal | 0.1145 | 0.105 | 0.1011 | 0.0962 | 0.0962 | 0.1036 | 0.0995 | 0.0999 | Uncal | Uncal | 0.1004 | 0.0932 | 0.1096 | 0.1037 | 0.1032 | 0.1038 | 0.1014 | 0.1043 | 0.1027 | 0.1025 | 0.0625 | 0.0623 | 0.1163 | 0.0978 | 0.0976 | 0.1011 | 0.0994 | 0.0998 | 0.0993 | 0.0933 | 0.0974 | 0.1076 | 0.1026 | 0.0794 | 0.1001 | 0.1028 | 0.0986 | 0.0945 | 0.0923 | Uncal | Uncal | 0.1042 | 0.1041 | 0.1029 | 0.1032 | 0.1 | 0.0996
SAMPLE, RPD Calculated | 2025NALS02293 A03 | -0.0003 u | -0.0007 u | 0.0043 | 0.0002 u | 0.003 | 0.0048 | 0.005 | 0.0221 | 0.022 | 0.0000 u | 0.0004 | Uncal | Uncal | 29.4233 | 28.9931 | 0.0002 | 0.0000 u | -0.0006 u | 0.0001 u | 0.0000 u | 0.0003 | 0.0002 u | 0.0001 | 0.2819 | 0.0004 | 0.0000 u | 0.0001 u | 0.0008 | -0.0001 u | 0.0000 u | 0.0095 | 0.0091 | -0.0001 u | -0.0008 u | 0.0014 | 0.0002 | 0.0005 | 0.0001 | 1.5117 | -1.5423 u | -0.0006 u | -0.0001 u | 1.6497 | 0.0035 | 0.0002 | Uncal | Uncal | 0 | 0.0005 | 8.8887 | 8.8506 | 0.0048 | 0.0063 | 0.0028 | 0.0033 | 5.4815 | 5.4966 | 0.0018 | -0.0101 u | -0.0007 u | 0.0006 | 0.0009 | 0.0000 u | Uncal | 0.0045 u | 0.0209 | -0.0001 u | 0.0001 u | 3.4935 | 3.5196 | -0.0002 u | 0.001 | Uncal | Uncal | -0.0034 u | 0.0027 u | 4.7145 o | 4.6166 o | 0.0000 u | 0.0007 | -0.0014 u | 0.0053 | 0.1317 | 0.1358 | 0.0009 | 0.0001 u | 0.0013 u | 0.0018 | 0.0039 | 0.0000 u | 0.0000 u | -0.0001 u | 0.0003 | -0.0072 u | 0.0001 u | -0.0002 u | 0.0000 u | 0.0609 | -0.0121 u | 0.0006 | 0.0208 | 0.0014 | -0.0013 u | Uncal | Uncal | 0.0001 | 0.0001 | 0.0077 | 0.006 | 0.0000 u | 0
SAMPLE DUPLICATE, RPD Calculated | 2025NALS02293 A04 | -0.0002 u | -0.0005 u | 0.005 | -0.0035 u | 0.0061 | 0.0048 | 0.0049 | 0.0221 | 0.022 | 0.0000 u | 0.0004 | Uncal | Uncal | 29.4265 | 28.8258 | 0.0001 | -0.0001 u | -0.0007 u | 0.0003 | 0.0002 | 0.0006 | -0.0020 u | 0.0002 | 0.2813 | 0.0003 | 0.0000 u | 0.0000 u | 0.0004 | -0.0001 u | 0 | 0.0108 | 0.0091 | 0.0000 u | -0.0005 u | 0.0012 | 0.0002 | 0.0005 | 0.0000 u | 1.0849 | 0.1124 u | -0.0003 u | -0.0002 u | 1.6463 | 0.0038 | 0.0000 u | Uncal | Uncal | 0 | 0.0006 | 8.9013 | 8.8665 | 0.0048 | 0.0061 | 0.0025 | 0.0026 | 5.4582 | 5.4574 | 0.0015 | 0.0094 | -0.0007 u | 0.0006 | 0.0012 | -0.0002 u | Uncal | 0.0239 | 0.017 | 0.0009 u | -0.0006 u | 3.4799 | 3.5118 | 0.0032 u | 0.0016 u | Uncal | Uncal | 0.0009 u | -0.0035 u | 4.7049 o | 4.6202 o | -0.0004 u | 0.0007 | -0.0013 u | 0.0052 | 0.1317 | 0.136 | 0.0008 | 0.0000 u | -0.0173 u | 0.0014 u | 0.004 | -0.0004 u | -0.0001 u | -0.0001 u | 0.0000 u | -0.0095 u | 0.0000 u | -0.0002 u | 0.0001 u | 0.06 | -0.0132 u | 0.0008 | 0.0163 | 0.0014 | -0.0007 u | Uncal | Uncal | 0.0001 | 0.0001 | 0.0075 | 0.0061 | 0.0000 u | 0.0000 u
LFM Background (10X Dil. of Sample) | 2025NALS02293 A03 - Dil. 10X | 0.0006 | 0.0001 u | 0.0007 u | -0.0028 u | -0.0008 u | 0.0005 | 0.0005 | 0.0021 | 0.0023 | 0.0000 u | -0.0003 u | Uncal | Uncal | 2.9521 | 2.8867 | 0.0001 | 0.0000 u | 0.0001 u | 0.0000 u | 0.0000 u | -0.0001 u | -0.0003 u | 0.0000 u | 0.0286 | 0.0002 | 0.0000 u | 0.0000 u | 0.0001 | -0.0001 u | 0.0000 u | 0.0016 | 0.0012 | 0.0000 u | -0.0007 u | -0.0005 u | 0.0001 | 0.0002 u | -0.0001 u | -0.7129 u | -1.7131 u | -0.0001 u | 0.0001 | 0.138 | 0.0002 | 0.0000 u | Uncal | Uncal | 0.0000 u | 0.0000 u | 0.8955 | 0.8873 | 0.0007 | 0.0014 | 0.0008 | -0.0001 u | 0.5415 | 0.5116 | 0.0003 | 0.0000 u | -0.0002 u | -0.0003 u | -0.0005 u | 0.0001 u | Uncal | -0.0004 u | 0.0064 | 0.0001 u | 0.0000 u | 0.3458 | 0.3524 | 0.0021 | 0.0002 u | Uncal | Uncal | -0.0085 u | -0.0078 u | 0.5022 | 0.4869 | -0.0003 u | 0.0002 | 0.0011 u | 0.0013 | 0.0132 | 0.0133 | 0.0001 u | -0.0006 u | 0.0070 u | 0.0008 u | 0.0000 u | 0.0003 u | 0.0001 | 0.0002 | -0.0024 u | -0.0052 u | 0.0074 | -0.0001 u | 0.0000 u | -0.0002 u | -0.0061 u | 0.0002 | 0.0002 | 0.0017 | -0.0008 u | Uncal | Uncal | 0 | 0 | 0.001 | 0.0009 | 0.0000 u | 0.0000 u
LFM Spike | 2025NALS02293 A05 | 0.0292 | 0.0225 | 0.1042 | 0.1002 | 0.0984 | 0.1037 | 0.1038 | 0.1053 | 0.1063 | 0.104 | 0.1069 | Uncal | Uncal | 3.1018 | 3.0629 | 0.102 | 0.1028 | 0.0922 | 0.0932 | 0.103 | 0.1026 | 0.1031 | 0.1033 | 0.13 | 0.1015 | 0.103 | 0.1026 | 0.1025 | 0.1049 | 0.1046 | 0.1039 | 0.1044 | 0.0993 | 0.1016 | 0.1018 | 0.1027 | 0.1003 | 0.0984 | 11.2645 | 9.8008 | 0.1016 | 0.1023 | 0.2502 | 0.1019 | 0.1025 | Uncal | Uncal | 0.1023 | 0.1025 | 1.0264 | 1.0154 | 0.1022 | 0.1043 | 0.1063 | 0.1063 | 0.6602 | 0.6262 | 0.0861 | 0.0792 | 0.0969 | 0.1029 | 0.1029 | 0.1022 | Uncal | 0.1224 | 0.1099 | 0.1009 | 0.0988 | 0.4368 | 0.4631 | 0.102 | 0.0991 | Uncal | Uncal | 0.0997 | 0.0993 | 0.6142 | 0.5885 | 0.103 | 0.1033 | 0.1033 | 0.1045 | 0.1159 | 0.1164 | 0.0627 | 0.0625 | 0.0975 | 0.0984 | 0.0972 | 0.1009 | 0.0994 | 0.1001 | 0.1013 | 0.0951 | 0.0998 | 0.1087 | 0.1029 | 0.0849 | 0.0999 | 0.1037 | 0.1029 | 0.0941 | 0.0937 | Uncal | Uncal | 0.1039 | 0.104 | 0.1033 | 0.1033 | 0.1001 | 0.0997
Instrument LFB | Lab Fortified Blank 1ppm May 29/25 ICV | 0.0321 | 0.026 | 0.094 | 0.0956 | 0.0948 | 0.0996 | 0.0997 | 0.1002 | 0.1017 | 0.1011 | 0.1013 | Uncal | Uncal | 0.0943 | 0.0954 | 0.1007 | 0.1001 | 0.0903 | 0.0909 | 0.1004 | 0.1003 | 0.0984 | 0.1004 | 0.0983 | 0.0977 | 0.1 | 0.0983 | 0.0993 | 0.1 | 0.1007 | 0.0988 | 0.1 | 0.096 | 0.0981 | 0.0997 | 0.0995 | 0.0975 | 0.0959 | 10.3399 | 8.4278 | 0.0986 | 0.0992 | 0.0972 | 0.0997 | 0.0999 | Uncal | Uncal | 0.1012 | 0.1006 | 0.0999 | 0.1 | 0.0986 | 0.0991 | 0.1028 | 0.103 | 0.0998 | 0.1001 | Uncal | Uncal | 0.0959 | 0.0996 | 0.1015 | 0.1006 | Uncal | 0.1006 | 0.1018 | 0.0987 | 0.0995 | 0.0909 | 0.0967 | 0.0954 | 0.0911 | Uncal | Uncal | 0.0951 | 0.0853 | 0.1 | 0.0987 | 0.0988 | 0.0997 | 0.0977 | 0.0993 | 0.0995 | 0.0991 | 0.0878 | 0.0874 | 0.1077 | 0.0993 | 0.0961 | 0.0975 | 0.098 | 0.0975 | 0.0949 | 0.1003 | 0.1027 | 0.1023 | 0.099 | 0.0761 | 0.0923 | 0.0996 | 0.0948 | 0.0927 | 0.0938 | Uncal | Uncal | 0.101 | 0.1012 | 0.1013 | 0.1005 | 0.0995 | 0.0968
Instrument Interference Check | Interference Check 1ppm May 29/25 ICV | 0.0306 | 0.0281 | 28.2122 o | 0.103 | 0.0844 | 0.0909 | 0.1858 | 0.1017 | 0.0997 | 0.0898 | 0.1001 | Uncal | Uncal | 0.1122 | 0.104 | 0.1011 | 0.1085 | 0.0791 | 0.0885 | 0.0216 | 0.1016 | 0.0969 | 0.098 | 0.1 | 0.0859 | 0.0962 | 0.0958 | 0.0965 | 0.0976 | 0.098 | 144.3939 o | 145.1832 o | 0.1035 | 0.1461 | 0.1025 | 0.1207 | 0.0959 | 0.0896 | 10.405 | 8.8638 | 0.0926 | 0.0953 | 0.1351 | 0.0958 | 0.0934 | Uncal | Uncal | 0.1028 | 0.0988 | 0.0983 | 0.0889 | 0.0969 | 0.197 | 0.0995 | 0.1023 | 0.1367 | 0.135 | Uncal | Uncal | 0.0925 | 0.1037 | 19.8724 o | 19.4050 o | Uncal | 0.1227 | 0.1054 | 0.0987 | 0.1169 | 0.0966 | 0.091 | 0.0986 | 0.0881 | Uncal | Uncal | 0.0779 | 0.0823 | 0.0919 | 0.0917 | 0.0991 | 0.0836 | 0.0978 | 0.0751 | 0.0992 | 0.0992 | 0.0755 | 0.0877 | 1.9872 o | 0.1164 | -0.0262 u | 0.2121 | 0.096 | 0.0982 | 0.093 | 0.0569 | 0.089 | 0.0972 | 0.1153 | 0.7374 | 0.3034 | 0.1044 | 0.0141 | 0.0921 | 0.0885 | Uncal | Uncal | 0.0993 | 0.099 | 0.1023 | 0.2285 | 0.0971 | 0.0993
Instrument ICV | ICV LVL 5 Jun 9/25 | 0.3987 | 0.3923 | 0.4965 | 0.4984 | 0.4988 | 0.4986 | 0.4989 | 0.5007 | 0.5044 | 0.5038 | 0.5023 | Uncal | Uncal | 0.498 | 0.498 | 0.5014 | 0.5002 | 0.4522 | 0.454 | 0.4973 | 0.4984 | 0.5062 | 0.5019 | 0.4966 | 0.4977 | 0.5021 | 0.4925 | 0.5024 | 0.5011 | 0.5038 | 0.5008 | 0.5034 | 0.4949 | 0.4984 | 0.5016 | 0.5007 | 0.4905 | 0.4852 | 47.8057 | 46.1059 | 0.5013 | 0.4999 | 0.5126 | 0.4994 | 0.4997 | Uncal | Uncal | 0.5055 | 0.5033 | 0.5039 | 0.5019 | 0.4963 | 0.4952 | 0.5202 | 0.5209 | 0.497 | 0.5054 | Uncal | Uncal | 0.4952 | 0.4962 | 0.5002 | 0.498 | Uncal | 0.519 | 0.4949 | 0.5 | 0.5134 | 0.4842 | 0.4889 | 0.4947 | 0.4965 | Uncal | Uncal | 0.496 | 0.5006 | 0.5171 | 0.5116 | 0.5015 | 0.4997 | 0.501 | 0.4964 | 0.5014 | 0.4989 | 0.4797 | 0.4819 | 0.4854 | 0.4956 | 0.4956 | 0.502 | 0.4849 | 0.4906 | 0.4954 | 0.4913 | 0.4984 | 0.5448 | 0.499 | 0.4778 | 0.496 | 0.5037 | 0.5152 | 0.4878 | 0.4868 | Uncal | Uncal | 0.5032 | 0.504 | 0.4992 | 0.5003 | 0.5026 | 0.4891
Digest Method Blank Comparison to Sample Matrix | DO NOT INPUT ANY DATA HERE
Digest Method Blank Comparison to 2.2X RL | DO NOT INPUT ANY DATA HERE
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
N/A | USED ONLY FOR ADDITIONAL QC
Reportable Limit (3xMDL) | Reportable Limit (No Dilution)
Reportable Limit (3xMDL) | Reportable Limit (No Dilution)
YELLOW IS PREFERRED WAVELENGTH
ICPOES (Reportable Numbers) *DO NOT ADJUST*
Sample Name | Comment
PERFORMANCE CHECK CALCULATIONS
ADJUSTMENT FOR <RL *DO NOT ADJUST*
Sample Name | Comment
PERFORMANCE CHECK CALCULATIONS
PERCENT RECOVERIES, RPD, or PASS / FAIL *DO NOT ADJUST*
Sample Name | Comment
*THIS TABLE HAS TWO ROWS REMOVED (IF EDITING, TAKE THIS INTO ACCOUNT)*

## SUMMARY TABLE VERIFICATION
REPORT ICPOES QA/QC DATA
| Total Metals | Units |  | IPC CCV5 after Calibration | IPC CCV5 after Batch | Digest Method Blank (LRB) | Digest LFB | SAMPLE DUPLICATE, RPD Calculated | LFM Spike | Instrument LFB | Instrument Interference Check | Instrument ICV | Digest Method Blank Comparison to Sample Matrix | Digest Method Blank Comparison to 2.2X RL
| Silver |  |  | 80.18 | 80 | PASS | 29.7 | PASS, Both <RL | 29.2 | 32.1 | 30.6 | 79.74 | N.A. | N.A.
| Aluminum |  |  | 101.74 | 100.84 | PASS | 103.2 | PASS, Both <RL | 104.2 | 94 | #VALUE! | 99.3 | N.A. | N.A.
| Arsenic |  |  | 99.54 | 99.3 | PASS | 96.7 | PASS, Both <RL | 100.2 | 95.6 | 103 | 99.68 | N.A. | N.A.
| Boron |  |  | 100.9 | 101.08 | PASS | 102.4 | 2.0202 | 103.8 | 99.7 | 185.8 | 99.78 | N.A. | N.A.
| Barium |  |  | 100.12 | 100.36 | PASS | 104.3 | 0 | 104 | 101.7 | 99.7 | 100.88 | N.A. | N.A.
| Beryllium |  |  | 103 | 103.32 | PASS | 105.8 | PASS, Both <RL | 106.9 | 101.3 | 100.1 | 100.46 | N.A. | N.A.
| Calcium |  |  | 99.84 | 99.7 | PASS | 97.4 | 0.010875 | Spike < 30% of Sample Matrix | 94.3 | 112.2 | 99.6 | N.A. | N.A.
| Cadmium |  |  | 99.96 | 100.06 | PASS | 102.1 | PASS, Both <RL | 102 | 100.7 | 101.1 | 100.28 | N.A. | N.A.
| Cerium |  |  | 100.22 | 100.28 | PASS | 92.4 | PASS, Both <RL | 92.2 | 90.3 | 79.1 | 90.44 | N.A. | N.A.
| Cobalt |  |  | 99.78 | 99.9 | PASS | 103.8 | PASS, Both <RL | 103 | 100.4 | 21.6 | 99.46 | N.A. | N.A.
| Chromium |  |  | 100.4 | 100.3 | PASS | 103 | PASS, Both <RL | 103.3 | 100.4 | 98 | 100.38 | N.A. | N.A.
| Copper |  |  | 100 | 100.1 | PASS | 103.2 | 0.21307 | 101.4 | 98.3 | 100 | 99.32 | N.A. | N.A.
| Dysprosium |  |  | 100.64 | 100.86 | PASS | 103.2 | PASS, Both <RL | 103 | 100 | 96.2 | 100.42 | N.A. | N.A.
| Erbium |  |  | 100.98 | 100.96 | PASS | 102.6 | PASS, Both <RL | 102.5 | 99.3 | 96.5 | 100.48 | N.A. | N.A.
| Europium |  |  | 102.06 | 102.16 | PASS | 104.7 | PASS, Both <RL | 104.6 | 100.7 | 98 | 100.76 | N.A. | N.A.
| Iron |  |  | 103.06 | 100.48 | PASS | 103.7 | 0 | 104.4 | 100 | #VALUE! | 100.68 | N.A. | N.A.
| Gallium |  |  | 99.88 | 100.08 | PASS | 98.3 | PASS, Both <RL | 99.3 | 96 | 103.5 | 98.98 | N.A. | N.A.
| Gadolinium |  |  | 100.58 | 100.78 | PASS | 102.4 | PASS, Both <RL | 102.7 | 99.5 | 120.7 | 100.14 | N.A. | N.A.
| Hafnium |  |  | 99.8 | 99.54 | PASS | 99 | PASS, Both <RL | 98.4 | 95.9 | 89.6 | 97.04 | N.A. | N.A.
| Mercury |  |  | 9580.7 | 9805.44 | PASS | 79.71 | PASS, Both <RL | 9800.8 | 84.278 | 88.638 | 9221.18 | N.A. | N.A.
| Holmium |  |  | 100.1 | 100.04 | PASS | 102.5 | PASS, Both <RL | 102.3 | 99.2 | 95.3 | 99.98 | N.A. | N.A.
| Potassium |  |  | 102.46 | 102.86 | PASS | 100.6 | 0.20631 | 112.2 | 97.2 | 135.1 | 102.52 | N.A. | N.A.
| Lanthanum |  |  | 101.1 | 100.64 | PASS | 102.7 | 8.2192 | 101.9 | 99.7 | 95.8 | 99.88 | N.A. | N.A.
| Lutetium |  |  | 100.12 | 100.28 | PASS | 102.4 | PASS, Both <RL | 102.3 | 101.2 | 102.8 | 101.1 | N.A. | N.A.
| Magnesium |  |  | 101.98 | 101.9 | PASS | 104.4 | 0.14165 | Spike < 30% of Sample Matrix | 99.9 | 98.3 | 100.78 | N.A. | N.A.
| Manganese |  |  | 100.82 | 100.68 | PASS | 102 | 0 | 101.5 | 98.6 | 96.9 | 99.26 | N.A. | N.A.
| Molybdenum |  |  | 100 | 100.32 | PASS | 105 | 11.3208 | 106.3 | 102.8 | 99.5 | 104.04 | N.A. | N.A.
| Sodium |  |  | 102.56 | 103.2 | FAIL | 102.3 | 0.71572 | Spike < 30% of Sample Matrix | 100.1 | 135 | 101.08 | <10% Matrix | >2.2X RL
| Niobium |  |  | 97.22 | 97.02 | PASS | 77.9 | PASS, Both <RL | 79.2 | #VALUE! | #VALUE! | #VALUE! | N.A. | N.A.
| Neodymium |  |  | 99.48 | 99.78 | PASS | 97.7 | PASS, Both <RL | 96.9 | 95.9 | 92.5 | 99.04 | N.A. | N.A.
| Nickel |  |  | 100.12 | 100.06 | PASS | 102.9 | PASS, Both <RL | 102.2 | 100.6 | #VALUE! | 99.6 | N.A. | N.A.
| Phosphorus |  |  | 102.6 | 101.36 | PASS | 105 | 20.5805 | 109.9 | 101.8 | 105.4 | 98.98 | N.A. | N.A.
| Lead |  |  | 99.58 | 99.38 | PASS | 101.1 | PASS, Both <RL | 100.9 | 98.7 | 98.7 | 100 | N.A. | N.A.
| Rubidium |  |  | 97.22 | 97.34 | PASS | 96.2 | PASS, Both <RL | 98.8 | 99.5 | 116.9 | 102.68 | N.A. | N.A.
| Sulfur |  |  | 98.72 | 100.5 | PASS | 103.6 | 0.22186 | Spike < 30% of Sample Matrix | 96.7 | 91 | 97.78 | N.A. | N.A.
| Antimony |  |  | 97.5 | 98.32 | PASS | 99.5 | PASS, Both <RL | 102 | 95.4 | 98.6 | 98.94 | N.A. | N.A.
| Selenium |  |  | 100.8 | 102.14 | PASS | 100.4 | PASS, Both <RL | 99.7 | 95.1 | 77.9 | 99.2 | N.A. | N.A.
| Silicon |  |  | 99.94 | 100.24 | FAIL | 100.4 | #VALUE! | Spike < 30% of Sample Matrix | 100 | 91.9 | 103.42 | #VALUE! | >2.2X RL
| Samarium |  |  | 101.64 | 102.18 | PASS | 103.2 | PASS, Both <RL | 103 | 98.8 | 99.1 | 100.3 | N.A. | N.A.
| Tin |  |  | 101.24 | 101.3 | PASS | 101.4 | PASS, Both <RL | 103.3 | 97.7 | 97.8 | 100.2 | N.A. | N.A.
| Strontium |  |  | 99.98 | 100.8 | PASS | 102.7 | 0 | 102.7 | 99.5 | 99.2 | 100.28 | N.A. | N.A.
| Tantalum |  |  | 101.88 | 101.36 | PASS | 62.3 | PASS, Both <RL | 62.5 | 87.4 | 87.7 | 96.38 | N.A. | N.A.
| Tellurium |  |  | 101.14 | 100.76 | PASS | 97.8 | PASS, Both <RL | 98.4 | 99.3 | 116.4 | 99.12 | N.A. | N.A.
| Thorium |  |  | 103.24 | 103.16 | PASS | 101.1 | PASS, Both <RL | 100.9 | 97.5 | 212.1 | 100.4 | N.A. | N.A.
| Titanium |  |  | 99.8 | 100.18 | PASS | 99.4 | PASS, Both <RL | 99.4 | 98 | 96 | 96.98 | N.A. | N.A.
| Thallium |  |  | 100.48 | 100.6 | PASS | 99.3 | PASS, Both <RL | 101.3 | 94.9 | 93 | 99.08 | N.A. | N.A.
| Thulium |  |  | 111.34 | 111.76 | PASS | 107.6 | PASS, Both <RL | 108.7 | 102.3 | 97.2 | 108.96 | N.A. | N.A. |  | Possible Comment Conditions
| Uranium |  |  | 100.64 | 100.58 | PASS | 100.1 | PASS, Both <RL | 99.9 | 92.3 | 303.4 | 99.2 | N.A. | N.A.
| Vanadium |  |  | 101.2 | 101.26 | PASS | 102.8 | 200 | 103.7 | 99.6 | 104.4 | 100.74 | N.A. | N.A. |  | a. IPCs are good, spike must be matrix relatated (SOP 8.1.1.3)
| Tungsten |  |  | 99.52 | 99.32 | PASS | 94.5 | PASS, Both <RL | 94.1 | 92.7 | 92.1 | 97.56 | N.A. | N.A. |  | b. Spike < 30% of Sample Matrix
| Ytterbium |  |  | 100.84 | 101 | PASS | 104.2 | PASS, Both <RL | 103.9 | 101 | 99.3 | 100.64 | N.A. | N.A. |  | c. Digest LRB <10% Matrix
| Zinc |  |  | 100.06 | 100.36 | PASS | 103.2 | 1.6529 | 102.4 | 100.5 | 228.5 | 100.06 | N.A. | N.A. |  | d. Digest LRB <2.2X RL
| Zirconium |  |  | 100.5 | 100.76 | PASS | 99.6 | PASS, Both <RL | 99.7 | 96.8 | 99.3 | 97.82 | N.A. | N.A. |  | e. Duplicate within 2.2X RL
f. IPC Blank within 2.2X RL
| Total Metals | Units | IPC Blank after Calibration (NALS Blank) | IPC CCV5 after Calibration | IPC CCV5 after Batch | Digest Method Blank (LRB) | Digest LFB | SAMPLE DUPLICATE, RPD Calculated | LFM Spike | Instrument LFB | Instrument Interference Check | Instrument ICV | Digest Method Blank Comparison to Sample Matrix | Digest Method Blank Comparison to 2.2X RL | Comments
IC Metal
Spike < 30% of Sample Matrix
Spike < 30% of Sample Matrix
Digest LRB <10% Matrix (Pass); Spike < 30% of Sample Matrix
Duplicate within 2.2X RL
Spike < 30% of Sample Matrix
LRB <10% Matrix, Spike < 30% of Sample Matrix
Duplicate within 2.2X RL
Digest LRB <2.2X RL (Pass)
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.